Name (print):_________________________________________   Job Title: ______________________________________________ Department/work area: ______________________________       Employee # ___________  NUID #_______Date ______________

#### Title of Competency: Roche Accu-Chek Inform II Glucometer

#### Job Category:  RN  LVN/LPN  MA/CNA/Tech Other

Rationale for Selection:  New  Change  High Priority  Problematic Regulatory Other __________________

Validator is expected to use guideline to provide consistency in validation process

| PERFORMANCE EXPECTATIONS | MSE | Met | Not Met* | Met with Remediation Revalidation |
| --- | --- | --- | --- | --- |
| PROCEDURAL STEPS |  |  |  |  |
| PATIENT TESTING |  |  |  |  |
| Checks patient for any indication that glucose testing may not be appropriate (See Procedure notes for patient assessment guidance). | V | x |  |  |
| Washes hands and don personal protective equipment (gloves, gowns, etc.). | V | x |  |  |
| Identifies the patient using two identifiers. | V | x |  |  |
| Explains the procedure to the patient. | V | x |  |  |
| Turns on the ACCU-CHEK Inform II meter. | RD | x |  |  |
| Enters their operator ID by means of barcode scanning or manual. NOTE: If the operator ID is not accepted, attempts to re-enter it. If it is still rejected, knows to contact their supervisor or Point of Care Coordinator. DOES NOT attempt to perform tests under another operator’s ID. | RD | x |  |  |
| From the Main Menu, touches Patient Test. | RD | x |  |  |
| Enters the patient identification in the ACCU-CHEK Inform II system by barcode scanning or manual entry. | RD | x |  |  |
| Confirms that the meter is coded (calibrated) to the same test strip code that is printed on the test strip vial by barcode scanning. Contacts their supervisor or Point-of-Care Coordinator if they are unable to confirm the correct test strip code. | RD | x |  |  |
| Removes a test strip from the vial and immediately recaps the vial. Inserts the test strip into the meter in the direction of the arrows and with the “ACCU-CHEK” lettering facing upward. | RD | x |  |  |
| Selects the appropriate puncture site and cleanses skin with alcohol pad. Allow puncture site to air dry completely. NOTE: If fingers are cold, warms fingers to increase blood flow and ensure adequate blood sample can be obtained. | RD | x |  |  |
| Punctures skin using an approved, single-use auto lancet. | V | x |  |  |
| Wipes away first drop of blood using gauze and applies second drop of blood to the front edge of the test strip. Does not squeeze finger to obtain blood sample and does not apply the blood sample to the top of the test strip. | V | x |  |  |
| Applies gentle pressure to the puncture site with a clean gauze square or cotton ball site for several minutes. | V | x |  |  |
| Identifies a critical blood glucose result. | V | x |  |  |
| Removes the test strip and dispose of it in the appropriate waste receptacle. | V | x |  |  |
| Touches to enter up to three appropriate comment(s), as appropriate. | RD | x |  |  |
| Touches the button to confirm the result and send the result from the meter wirelessly (This functionality may not be active yet). | RD | x |  |  |
| Cleans and disinfects meter. | V | x |  |  |
| Returns meter to base unit for charging. | V | x |  |  |
| Documents the result in the patient chart in KP HealthConnect (This step may not be required if results are wirelessly). | V | x |  |  |
| States appropriate actions for follow up on results that exceed critical or reportable limits according to policy. | V | x |  |  |
| QUALITY CONTROL TESTING |
| Turns on the ACCU-CHEK Inform II meter. | RD | x |  |  |
| Enters operator ID by means of barcode scanning or manual entry. | RD | x |  |  |
| From the Main Menu, touches Control Test. | RD | x |  |  |
| Selects the level of control that you wish to test: Level 1 (Lo) or Level 2 (Hi). | RD | x |  |  |
| Enters and confirms the lot number of the control solution by means of barcode scanning or manual entry. | RD | x |  |  |
| Confirms that the meter is coded (calibrated) to the same test strip code that is printed on the test strip vial by barcode scanning. | RD | x |  |  |
| Removes a test strip from the vial and immediately recaps the vial. Inserts the test strip into the meter in the direction of the arrows and with the “ACCU-CHEK” lettering facing upward. | RD | x |  |  |
| Applies control solution to the front edge of the test strip. | RD | x |  |  |
| States if the control test has passed or failed and explains the appropriate actions to take if the test failed. | V | x |  |  |
| Removes the test strip and disposes of it in the appropriate waste receptacle. | RD | x |  |  |
| Touches the comment button ( ) to enter an appropriate comment(s). | RD | x |  |  |
| Touches the button to confirm the result and send the result from the meter wirelessly. | RD | x |  |  |
| Returns the meter to the base unit for charging. | V | x |  |  |
| METER CLEANING |
| States when meter should be cleaned. | V | x |  |  |
| Turns the meter off. | V | x |  |  |
| Squeezes excess cleaning liquid from disinfectant wipe. | V | x |  |  |
| Demonstrates proper cleaning procedure: Wipes face of meter surface 3 times horizontally Wipes face of meter surface 3 times vertically Wipes sides of meter Wipes back of the meter Ensures that no liquid enters the test strip port | V | x |  |  |
| Verbalizes appropriate contact or dwell time (the duration of time required for the surface to remain wet for proper disinfection) based on type of cleaning wipe used. | V | x |  |  |
| Allows meter to dry completely prior to docking on the base unit. | V | x |  |  |
| CRITICAL THINKING (LVN/LPN/MA/other – Problem solving question – See Guideline) |  |  |  |  |
| What action do you take if the meter is not coded to the same test strip code that is printed on the test strip vial? | V | x |  |  |
| What action do you take if the meter does not turn on? | V | x |  |  |
| INTERPERSONAL SKILL (Diversity, Cultural, Age Specific, Customer Service – See Guideline) |  |  |  |  |
| What do you tell a patient if the blood glucose result is “HI’? | V | x |  |  |
| What do you do if a patient requests to use an alternate testing site (i.e. the palm of the hand, forearm, etc.)? | V | x |  |  |

Method of Skill Evaluation (MSE) - O (Observation) RD (Return Demonstration) WA (Written Assessment) V (Verbalization) CR (Chart Review)

References:

CLSI Document H4-A5: Procedures and Devices for the Collection of Diagnostic Blood Specimens by Skin Puncture – Approved Standard; Fifth Edition; 2004; Page 11

CLSI Document POCT12-A3: Point-of-Care Blood Glucose Testing in Acute and Chronic Care Facilities; Approved Guideline – Third Edition; 2013; Section 10.2.3; Page 15

Lippincott Procedures. (2020). Point-of-Care Testing, Ambulatory Care. Retrieved from

https://procedures.lww.com/lnp/view.do?pId=3326003&hits=glucose,blood,testing,test,tests&a=false&ad=false

Roche Diagnostics. (2017). Accu-chek Inform II Operator’s Manual

Validated by: (print name/signature/dept.)___________________________________________________ Date ___________ Employee: (print name/signature/dept.)______________________________________________________ Date ___________ Outcome: *needs to complete plan & then revalidate satisfactorily completed competency able to validate others Plan required for “Not Met” - Notify manager if employee judged not competent